# Physiological Data Collection Protocol: Noise & Artifact Classification

Date: November 2025

Investigator: Wearable and Biosensing lab (WBL), University of rhode island.

Total Duration: 20 Minutes

Device: BIOPAC Systems

## 1. Objective

To collect a physiological dataset (ECG/PPG) containing strictly labeled intervals of valid physiological signals and specific noise artifacts. This data will serve as "Ground Truth" for evaluating the performance of unsupervised clustering algorithms.

## 2. Experimental Design

- Valid Signal Types: Resting Baseline (Static) and Deep Breathing (Dynamic).

- Noise Types: Motion Artifacts (Phone Scrolling), EMG Noise (Muscle Stiffness), and Sensor Failure (Loose Contact/Disconnect).

- Block Design: Alternating 1-minute blocks of activity and  2 minutes for recovery.

## 3. Equipment & Setup

- Software: AcqKnowledge (BIOPAC).

- Sampling Rate: Recommended 1000 Hz.

## 4. The 20-Minute Protocol Schedule

| Time Window | Class | Activity | Detailed Instructions for Participant | Label ID |
| --- | --- | --- | --- | --- |
| 00:00 – 02:00 | Clean | Resting Baseline | Sit comfortably in a chair. Hands resting on lap. Breathe normally. Do not speak or move. | 0 |
| 02:00 – 03:00 | Noise | Phone Scrolling | Hold a mobile phone in the hand with the sensor (or dominant hand). Scroll through social media reels continuously. Use normal thumb movements. | 1 |
| 03:00 – 05:00 | Clean | Deep Breathing | Put the phone down. Do not move body. Inhale deeply for 4 seconds, exhale for 4 seconds. Focus on rhythmic chest expansion. | 0 |
| 05:00 – 06:00 | Noise | Body Stiffness | Make the entire body rigid. Clench fists, tighten biceps, and stiffen the torso. Hold tension while breathing shallowly. | 2 |
| 06:00 – 8:00 | Clean | Resting Recovery | Release all tension immediately. Exhale sharply. Return to normal, shallow breathing. Sit perfectly still. | 0 |
| 8:00 – 9:00 | Noise | Sensor Loose | Do not unplug. Gently wiggle the sensor clip or peel the electrode tape slightly so it makes poor contact. Aim for chaotic spikes on the screen. | 3 |
| 9:00 – 11:00 | Clean | Deep Breathing | Secure the sensor back in place. Resume deep, rhythmic breathing (Inhale 4s / Exhale 4s). | 0 |
| 11:00 – 12:00 | Noise | Chair Restlessness | Shift weight from side to side, cross and uncross legs, lean forward and back, and adjust posture frequently. | 1 |
| 12:00 – 14:00 | Clean | Pre-Fail Baseline | Sit perfectly still. Ensure the signal has stabilized and is clean. | 0 |
| 14:00 – 15:00 | Noise | Full Disconnect | Unplug the sensor lead completely. Remove the clip from the finger or unsnap the lead. Drop the wire on the table. | 3 |

## 5. Data Labeling Reference (Ground Truth)

Use the following Integer IDs when creating the evaluation matrix (Confusion Matrix) for your clustering algorithm.

- Label 0 (Valid Signal): Includes both "Resting" and "Deep Breathing." The algorithm must learn that rhythmic respiratory wander is not noise.

- Label 1 (Motion): Low-frequency baseline wander caused by phone usage.

- Label 2 (EMG): High-frequency "fuzz" caused by muscle tension.

- Label 3 (Failure): Signal saturation, rail, or flatline caused by loose contact or disconnection.